## Revision Control
Revision Control
Revision | Revised by | Revision date | Change Description
0.1 | Maja Husum | 2024-04-04 | First Version of risk analysis

## Usability Risk
Hazard ID# | Hazard | Cause of hazardous situation | Sequence of events | Hazardous situation | Harm | Object at risk |  |  | Risk estimate |  |  | Risk Mitigation | Residual risk estimate |  |  | Is risk acceptable? | New hazards or hazardous situations?
[6 empty] | Patient | Operator | Equipment | O1 | S1 | RF1 |  | O2 | S2 | RF2
UR1 | Critical performance | User enters erroneous data in the form of date | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 3 | 2 | 6
UR2 | Critical performance | User enters erroneous data in the form of timestamp | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 4 | 2 | 8
UR3 | Critical performance | User enters erroneous data in the form of daily evaluation | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 2 | 2 | 4
UR4 | Critical performance | User enters erroneous data in the form of tracking accident or no accident | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 4 | 2 | 8
UR5 | Critical performance | User misinterprets the functionality of the daily evaluation | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 2 | 2 | 4
UR6 | Critical performance | User enters erroneous data for previous days, in the form of daily evaluation, due to missing daily evaluation | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 4 | 2 | 8 | hvis ikke du kan huske dagen, så lad være
UR7 | Critical performance | User enters erroneous data for previous timestamp, in the form of tracking accident or no accident | 1. Erroneous data is visualized 2. The user interprets the visualized data 3. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 5 | 2 | 10
UR8 | Data access | User is technically incompetent | 1. The user is unable to connect Ucon to their mobile phone | Data from UCon is unavailable | Negligible | x |  |  | 3 | 1 | 3
UR9 | Alarm | User turns notifications off | 1. The user is not receiving notifications 2. The user is not reminded to use the system | The user aborts the use of the system | Negligible | x |  |  | 3 | 1 | 3
UR10 | Alarm | User enters erroneous user defined notification settings | 1. The user receives an erroneous notification | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 3 | 2 | 6
UR11 | Data access | User enters invalid login information | 1. The user is unable to use the system | Data from UCon is unavailable | Negligible | x |  |  | 4 | 1 | 4

## Software Risk
Hazard ID# | Hazard | Cause of hazardous situation | Sequence of events | Hazardous situation | Harm | Object at risk |  |  | Risk estimate |  |  | Risk Mitigation | Residual risk estimate |  |  | Is risk acceptable? | New hazards or hazardous situations?
[6 empty] | Patient | Operator | Equipment | O1 | S1 | RF1 |  | O2 | S2 | RF2
SWR1 | Alarm | The system sends the user too many notifications | 1. The user has turned all notifications on | User is over-stimulated | Negligible | x |  |  | 4 | 1 | 4
SWR2 | Alarm | The system is incompatible with the operating system | 1. The system sends the user illegible notifications | User is not informed | Negligible | x |  |  | 4 | 1 | 4
SWR3 | Alarm | The system sends the user unwanted notifications | 1. The user has turned all notifications off 2. The system does not register the disabled notifications | User is misinformed | Negligible | x |  |  | 2 | 1 | 2
SWR4 | Alarm | The system does not send the user notifications | 1. The user has turned all notifications on 2. The system does not register the enabled notifications | User is not informed | Erroneous treatment | x |  |  | 2 | 2 | 4
SWR5 | Data integrity | The system is incompatible with the users mobile phone | 1. The system retrieves an incorrect date 2. The system displays an incorrect date | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 2 | 2 | 4
SWR6 | Data integrity | The database saves a daily evaluation entry incorrectly | 1. The system displays incorrectly saved data 2. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 1 | 2 | 2
SWR7 | Data integrity | The database overwrites previous daily evaluations | 1. The system displays incorrectly saved data 2. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 1 | 2 | 2
SWR8 | Data access | The system is unable to connect to the database | 1. The system can not visualize data | Data from the database is unavailable | Negligible | x |  |  | 2 | 1 | 2
SWR9 | Data availability | There is no available storage on the mobile phone for the database | 1. The system cannot store entered data | The system is unusable | Negligible | x |  |  | 1 | 1 | 1
SWR10 | Data access | The bluetooth connection is attacked by a malicious source | 1. The malicious source can monitor stimulation data | Sensitive data is accessed | Data breach | x |  |  | 1 | 1 | 1
SWR11 | Data access | The system is unable to connect to Ucon | 1. The Ucon is out of range | Data from UCon is unavailable | Negligible | x |  |  | 2 | 1 | 2
SWR12 | Data access | The system is unable to connect to Ucon | 1. The Ucon is turned off | Data from UCon is unavailable | Negligible | x |  |  | 4 | 1 | 4
SWR13 | Data integrity | The system erroneously reads services from Ucon | 1. The system visualizes data incorrectly 2. The user acts on the basis of the visualized data | The user receives erroneous stimulation | Erroneous treatment | x |  |  | 1 | 2 | 2
SWR14 | Data integrity | The system erroneously writes services to Ucon | 1. The system receives unexpected data | The system does not receive requested information | Negligible | x |  |  | 1 | 1 | 1
SWR15 | Data transfer | Packet loss in the Bluetooth connection | 1. The system query data from Ucon 2. Data is lost | The system does not receive requested data | Negligible | x |  |  | 2 | 1 | 2
SWR16 | [10 empty] | 0
SWR17 | [10 empty] | 0
SWR18 | [10 empty] | 0
SWR19 | [10 empty] | 0
SWR20 | [10 empty] | 0
SWR21 | [10 empty] | 0
SWR22 | [10 empty] | 0
SWR23 | [10 empty] | 0
SWR24 | [10 empty] | 0
SWR25 | [10 empty] | 0
SWR26 | [10 empty] | 0
SWR27 | [10 empty] | 0
SWR28 | [10 empty] | 0
SWR29 | [10 empty] | 0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0